Service Report

| Company: | Novartis Basel (Gaby Rummel) XYZV67 |
| --- | --- |
| Date: | 29 January 2007 |
| Service Engineer: | JS |
| Details: | Went there to replace a motor on a USD, but found that actually the USD leadscrew had seized. After taking it apart and loosening the tension on the far bearing it started moving again. Ran it in for a few cycles using debubble. Installed new software V 22.06.0190 |

Signed: