Service Report

| Company: | ICR – Chester Beaty |
| --- | --- |
| Date: | 21.7.06 |
| Service Engineer: | Mike Stevens |
| Details: | The plateloader removed from the box, the arms inserted. Then connected to a DB5 MCC. The Z & V axis both ran at 1.6 time’s normal speed, (30*1.6) without problem. Patrick noticed that it only took a very small amount of pressure to make the arms miss steps at this speed. When system was ran at normal speed however it took substantially more pressure and still did not miss steps. The motors were replaced with the new, more powerful motors. For rezero stops the new leadscrews have collars with grub screws that are tightened onto a flat on the leadscrew. Upon testing, both axis performed perfectly both in axis test and rezero. The speed was increased from 30.00 to 52.00, as is used for these motors in the current system. Up to speeds of 1.4 both axis operated correctly. At 1.5 both axis started to miss steps. The entire system was recalibrated, new files produced and a new disk burnt. A new plateloader lead was also taken in case the problems have been due to a worn plug / socket causing bad connections. |

Signed: