Service Report

| Company: | ICR |
| --- | --- |
| Date: | 26.7.06 |
| Service Engineer: | Mike Stevens |
| Details: | ICR was visited by myself on the 25.7.06. The old software was removed, the new software was installed (Version V18.75.0433) and the machine tested. When a vapour diffusion experiment was run, both the tip and the Gilson tip hit the edge of a bottle in the rinse block. The rinse block was checked and found to be “bridging” slightly in the centre, but by too small an amount to cause the collision to occur. The block was gently bent back on itself; this reduced the bridging further still. A new cable was installed, and the experiment repeated. The collision occurred again. Upon consulting Peter it became apparent that although the tip height had been set to the plate at the start of the experiment it had not been set initially using the “Install Tip” command from the front panel menu. This was the reason for the collisions and there was no fault with robot. Staff at ICR assert that they had not been show this feature of the system either when it was installed or at any time since. |

Signed: