Service Report

| Company: | University of Michigan |
| --- | --- |
| Date: | 4/4/06 |
| Service Engineer: | M. Stevens, J. Smith |
| Details: | When this MCC was returned to us it would not start up. Upon opening the case, damage was found on inside of the right hand side panel and to several of the larger capacitors on the motherboard. This indicates that the heatsink / fan assembly for the CPU had been knocked loose by a particularly harsh knock during transit and had bounced around inside the case. Upon removing the heatsink / fan assembly the CPU chip was examined and found to be severely damaged by overheating. A new CPU chip was tried but the system still would not boot up. This indicates damage to the motherboard. Note: There was no heat sink paste found upon the CPU chip. This is contrary to the report from Michigan, that stated a computer technician had used heat sink paste when he replaced the heat sink / fan assembly. The Flash drive and cards were transferred to a new case containing a new motherboard, CPU and fan assembly. Bios was set and the system was switched on and allowed to run overnight. CPU temperature stayed at a steady 37°C. MCC attached to a XYZV and a rezero of all axis performed. MCC operates correctly. Michigan advised to check heat sink and fan are in place before starting up machine. Heat transfer paste provided and instruction on correct use sent with MCC. |

Signed: