Novartis Service Visit 14 March 2006

Send animated GIF to Allan

Send new CD for seeding

Send quote for ASSO

Send notebooks (also to customers of CCs)

Look up NIH meeting

Send samples of Duos

Grease needed in tool case

Blue tips needed in tool case

A little bit of emery in tools

Big syringe in tools

Spare 2-bore collet

- Allan’s number 41 61 696 20 33.  See card for address.

- Allan and Fred were confused by the Show Me aspect of the WASPRUN interface.

- Set up experiment 5.15  – 0.3 + 0.3 ul to Droplet2.  No protein or precipitant were dispensed.  Log shows 0 + 0 ul.

- 0.3 + 0.3 to Droplet - OK

- Traced in checking only, stepwise, Intelliplate.  At line 577, Multdrop was 0, yet did Transfer (vStock etc  (This was expected in first pass probably)

- Also didn’t work with Greiner Flat

- Droplet and Droplet2 – worked fine.

- James’ special seeding CD has disintegrated.

- Same problem showed up on my Notebook.  However, it was working fine at Novartis before installing Seeding script.

- Peter and I saw no dispensing to Droplet2.  This was due to missing line, Drops = Drops + 1

- Allan saw tip going to Droplet instead of Droplet2.  On inserting the missing line, I now see Allan’s problem(!).  WASP says it is dispensing to Droplet 2, but actually dispenses to Droplet.  This is due to old software and new script.

- Allan had 18.69.  I am going to install 18.72.  This did not appear to change the MCC settings.

- I now realize that Allan had rather older software than I thought.  He obviously asked me to leave it alone last time.  I will now install 18.73, and try screening software.

- 18.73.0127 has