Service Report

| Company: | Nerviano Medical Sciences Srl |
| --- | --- |
| Date: | 4-02-2008 |
| Service Engineer: | James Smith |
| Details: | Robot tested and calibrated. Test Results: X-axis runs at 1.8x Y-Axis runs at 2.0x Z-axis runs at 2.0x V-axis runs at 1.7x Minimum required: 1.4x USDs run through full range All tests passed New software installed: Version 23.49.0720 with modification to Wasp to run on Windows 2000. Note: Newer versions of our software will not support Windows 200 in future. It is recommended that the PC be upgraded to Windows XP or Vista with minimum 512 MB of RAM before the next service visit. User training provided on new experiments (v8.3) Follow up: A wide rinse-block with protein loading strip will be sent to the customer for best use of new experiments. |

Signed: