Service Report

| Company: | University of Michigan – Oryx4 |
| --- | --- |
| Date: | 22 February 2007 |
| Service Engineer: | James |
| Details: | Service contract visit as per request. Installed new software Version 22.08.0320 and trained users in its use. Checked machine for general condition and found it to be in good mechanical condition. Created gradient seeding(Additive) experiment for Microbatch Changed Oil tip as the old one was damaged. Improved labelling on the Oil bottle to make it clear that pure paraffin should be used for MB screening and optimization. |

Signed: