### UGA – service visit.

September 22, 2005

UGA II

- Win 2000

- V arm was reported to be going too low during runs.

- Software errors at start.  (Log showed last expt 2 Sept).  (1) menu on FP greyed out and inactive.  FP says closed.  Clicking on Terminal – Term jumps to front then FP jumps in front of it.   Can’t get focus on Term.

- V at 1.0 OK.  V1.2 OK.  V1.4 OK. V1.5 OK

- Therefore assumed this was a software problem.

- Log – they tried about 20 times to run expt.

- Old software: ran MB_trans.  After 23 wells V tip too low.  Above software errors also present.

- Ran MB_trans.  After 12 wells V tip too low.  No software errors.  V at 1.5 OK

- Installed new software.  Changed file to do double screen scripts (large syr on Channel 5).

- Ran MB_trans.  OK

- Ran three VD (CC) expts.  OK

- X at 1.3 OK.  Y at 1.3 OK.  Z at 1.3 OK.  V at 1.3 OK

- Changed XY offsets for table slightly

- RSDUMP – 1 gross error on X

- Rinse block bowed up in middle.

UGA I

- Rinse block bowed up at ends.  Straightened.

- X1.3 OK, Y 1.3 OK, Z 1.7 OK, V 1.3 OK

- Changed speed of Z from 16 to 20.  Z 1.3 still OK.

- Bret moved UGA I to anaerobic chamber

Bret Dillard  eandb@uga.edu

Old version of VD cards !!!!  Send new one * 2, + 1 for Renu.

Send evaporation shield.

UGA III (Renu)

- X Y Z V all OK at 1.3

- Aligned arm with hole in frame.  Miles out.  Undid Z-box and moved closer to hold.  Ended up about 1.5 mm from hole.