# Beth Israel Deaconess 19 Feb 2007

617 754 1223 Xiang Xu

617 754 1224 Lin Lin

617 754 1257 Bench

617 754 1203 Ming Dong

James

- One well of Duos is scratched.

- Compared to –P plate the VB plate seemed to be badly defined.  The VB plate was too high by one big click.

Peter

Info

- New software (v21) was not installed.

- Collet with rubber extension had not been taken out of packet.

- They don’t have rinse block with protein strips.

- FindIndexP – no gross or small errors

- Xoffset was 2.11 Yoffset -0.84 on dowel hole before pressing any keys.  We received a warning.  I set those values in 2006.  Changed to 0.59, -0.05 now.

- You have to Go To Z = -5 to get to tip to touch table.  Goes down very slowly.  This was due to tip being wrongly adjusted, presumably because of Bug in previous software (v18).  Or could have been using the wrong plate (-d / -p).

- 11.12 am – missed lots of steps on X or Y.  Because I had tightened up Fork

- Syringe 6 jammed near bottom.  Too dry?   Cut off a few threads from both ends.  Ran-in successfully at 8.0

- S1 same problem.  Cut off threads here too.

- Took logs file.  Must rename log folder on silver memory stick.

- Moved S5 to S6.  Took away original S6 with 100 ul syringe clamp from S5 and thumbscrews.  Send back.  Also send back Rinse Block with protein strip positions.