Service Report

| Company: | # University of Michigan |
| --- | --- |
| Date: | 31/7/06 |
| Service Engineer: | M. Stevens |
| Details: | The plateloader was examined and found to be very stiff in the X-axis. The belt was slackened and the motor and pulleys were found to run smoothly. The table itself was stiff. It was examined and found that the cause of the problem was that some kind of viscous material, of unknown origin, had got onto the front shaft, from where it had transferred and congealed on the lower button, in the fork bearing assembly that supports the front of the table. This “gummed” the bearing up, and possibly being absorbed into the plastic button, causing it to swell and thus clamp down on the shaft. Both buttons were replaced and this has cured the problem. The table was recalibrated in the X & Y axis. The readings were so close to the original that upon consulting with James it has been decided that new software need not be issued. |

Signed: