Service Report

| Company: | Novartis Basel (Allan) XYZV-50 |
| --- | --- |
| Date: | 29 January 2007 |
| Service Engineer: | JS |
| Details: | Replaced all long belts. Old belts were discoloured probably due to UV and this can be clearly seen as both belts are discoloured on one half only (that which is closest to the inside of the room and to the left.) Replaced left-hand nearside corner pillar with new pillar with molybdenum button as old Teflon button was very warn. Reinstalled new software V 22.06.0190 – issues with DAO350.DLL not being installed into C:\Program Files\Common Files\Microsoft Shared\DAO this causes FP and XS to crash on startup. Suggestions: Allan should send his plateloader back for a service when he can spare it for a week or two. We should replace the X and Y belts and pulleys with the 2mm black belt version as these are supposed to be more resistant to UV. |

Signed: