Service Report

| Company: | Uppsala |
| --- | --- |
| Date: | 27-03-07 |
| Service Engineer: | James |
| Details: | Installed CD 22.12.0330 Error- Uppsala hardware folder did not contain the channels file for Oryx6[on8] 2x 100 ul syringes were broken and others showed signs of being run up to the top. I think rezeros had been done without waiting for all syringes to get to the bottom. Trained more users on Optimization and the new software. |

Signed: