# Low-level Commands for MCC

### Forth Words present in MCC:

FindIndexPositions 	– calibrate X and Y with respect to limit switches

rsdump 		– lists errors for X and Y

mdump 		– lists parameters of all motors

sdump 			- lists step situation for all motors

bitmask actuate 	- where bitmask is as follows

BCOMP           		=       0x0001	=1

VECTOR          		=       0x0002	=2

SPEED           		=       0x0004	=4

CONCURRENT      	=       0x0008	=8

LIMIT_SWITCH    	=       0x0010  	=16

BUTTON          	=       0x0020	=32

KEYBD           		=       0x0040	=64

LIMITS          		=       0x0080	=128	Won't allow movement beyond limits - not limit switch

NO_ACCELERATION =       0x0100	=256

LS_HIGH         		=       0x0200	=512

NO_WAIT        	=       0x1000	=4096

DEFAULT_MODE    =       SPEED+BUTTON

LIMIT_SWITCH    		=   0x0010  	=16	- move until opto beam broken

LIMIT_SWITCH+LS_HIGH	=  0x0210  	=528	- move until opto beam transmitted

Use e.g. speed, button, limit switch, LS_high, vector, concurrent =

Moves:

Real:		Mno:int Distance:float {moveBy | moveTo} eg 10 50.0 moveTo 46 actuate

Steps:		Mno:int Steps:Long     {stepBy | stepTo} eg 10 200L stepBy 46 actuate

(46 actuate is suitable for the PlateLoader because it does a dog-leg)

(Mno:int means that the motor number is an integer.)

{} shows alternatives.

axisX 50.0 moveBy

Xposition:float Yposition:float XYGoto

Speed:

Mno:int Speed:float SetSpeed   /   Mno:int GetSpeed (f. – this prints a floating point no.; “.” prints an integer)

Mno:int SpeedFactor:float SetSpeedFactor	(ssf Mno speed  - at keyboard )

Read Limit Switch:

LS Mno (gives 0,1,2,3 = 2-bit pattern for up to two limit switches on each motor)

To do TestAxes with just one motor:

1 FindVMax to do TestAxes on just e.g. motor 1

### PlateLoader Test Words   -  type <test> first

fi        		{ does FindIndexPositions

rall      		{ rezeros all axes

.LS       		{ display all limit switch values continuously, until a key is hit

TestAxes  	{ finds maximum reliable speed for each axes

### One motor at a time - for efficiency

Axis numbers are X, Y, Z, V = (0, 1, 2, 3)

<b> refers to an axis bitmap; 2^X=1, 2^Y=2, 2^Z=44, 2^V=8; so {X|Y|Z} = 7

<b> _fip  { FindIndexPositions for one or more motors in <b>

<A> FindVMax 	{ Find maximum speed of axis <A>             eg Z FindVMax

### PlateLoader Quick Move Words   -  type <test> first

fi        		{ does FindIndexPositions

rall      		{ rezeros all axes

.LS       		{ display all limit switch values continuously, until a key is hit

x y z v   	{ Sets X, Y Z or V axis as target. Subsequent commands use this axis

(mm) >    	{ Sets target position for current axis

(mm) +>   	{ Sets relative target position for current axis

(mm) >>   	{ Sets target and moves

a         		{ Cause any pending moves to happen.

jog       		{ go into JOG mode. Use <+>/<->, and <Enter> to quit

(v) speed 	{ Sets the speed to (v)

Example:   x 60 > y 75 >>     will move to (X, Y) = (60.0, 75.0)

52 speed z 50 > v 40 > a   will raise Z and V axes at full speed.

### PlateLoader Calibration

Type <test> first. Then the following words become available

fi    		{ does FindIndexPositions

rall  		{ rezeros all axes

.LS   		{ display all limit switch values continuously, until a key is hit

CalX  		{ calibrate the X axis

CalY  		{ calibrate the Y axis

CalZ  		{ calibrate the X axis

CalV  		{ calibrate the Y axis (only on XYZV, not Nano)

CalX/Y/Z/V produce calibration files called XCAL.DAT, YCAL.DAT etc.

Front Panel has a menu item {PlateLoader | Transfer Calibration Files} to

download them to "C:\Program Files\Douglas Instruments\MCC" on the PC.

### USD Instructions (BallScrew USD2.0)

Backlash measurements are done by comparing 10 readings of the unloaded USD

with 10 readings of the USD loaded with a 100ul syringe. The unloaded readings

are taken with a rubber band providing light tension, so that backlash is

prevented. This therefore measures the hysteresis of the optoswitch.

The optoswitch must be connected to channel 10 (Y axis LS input)

### USD Test Words        - type <test> first

sno b   		} Does 10 hysteresis measurements, followed by 10 backlash measurements

} and reports the means and the difference, which is the backlash.

sno s+  		} moves syringe sno +0.05mm   eg   2 s+

sno s-  		} moves syringe sno -0.05 mm

sno s++ 	} moves syringe sno +1.0 mm

sno s-- 		} moves syringe sno -1.0 mm