Saskatchewan Uni Service Visit RPT April 10 2006

- Remember to put strip of Emery into tool kit

- Replenish special plastic grease

- Put in stainless adjustment frame.

- Send samples of Duo.

- Send new 2-bore collet.

- Bringing back USDs.  (Send back rezeroed.)

- Send spare 5 ml glass syringe

- Upgraded MCC Feb 04 to Dec 05

- Upgrade Flash Firmware – to transfer Hardware.fth only.  Didn’t work – Terminal crashed.  Host to MCC worked.  Upgrade FF crashed again.

- S2 not moving.  V stiff.  Slowed to 1.5.  Still not moving.

- S1 – S5 all are not working.

- V axis judders up and down when you rezero it.**

- Came back.  Lights 5 6 X Y Z V on MCC were lit.  Closed FP – OK

- Could not turn S5 more than 1/8 turn with fingers.  Took to pieces.  Only 1 grubscrew was done up on dog.  Plastic nut came out.  Put it in again with loctite.

- Rezeroed / debubbled all motors, still at 1.5.  S5 Missed steps intermittently.  S1-S4 didn’t move.

- Pitches seemed to be wrong.  Tried cleaning with acetone and lubricating with RS plastic grease. No better.

- Run-in did not work.  After about 2 hours I found out that the speeds were being set to 0.  Altered file to hardwire speed.  Started 1.0.  WHAT FILE DO WE USE?  WHERE IS IT?

- All nuts could easily be removed by hand, or came out spontaneously.  Trimmed both ends with a very pointed sharp scalpel.  Had to repeat 3 times on some nuts.  Punched them in place in 3 places with a Posidrive bit.

- ** Took ZV box to pieces.  All looked OK.  Resoldered one connection on PCB to D-plug, but it was probably OK anyway.  Took out the 4 crimped joints of the V connector on the top of the PCB and applied solder.  Still judders.  Tried slower – fine.  Slowed up to 45 from 55

- Ran in all syringe drives overnight (450 reps) and another 200 at high speed next day.  USDs are numbered.

- I always have to file the edge of the evap shield – too obtuse angle.

- V stop popped off.  Put on again with Green Loctite.

- Final table offsets: 3.29, -0.13, (-2.2)

- Z was far too high.  Adjusted down with grubscrew about 2 mm.

- Far too many litter pickers.  Removed all except 1.

- How are you supposed to use the protein salt in MB to VD converter?  You can’t easily increase/decrease the protein salt in the MB/VD protein.  E-mail Jenny – puttick@sask.usask.ca

- Lost 2-bore collet.  It was rough anyway.